Kepco SN488 Power Supply Programmer

________________________________________________________________________________

Introduction:

This instrument module provides GPIB programming support for the

Kepco SN488 Power Supply Programmer.  It contains functions for

opening, configuring, setting output, and closing the instrument

via the GPIB

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

- the instrument is connected to the GPIB.

- the GPIB address supplied to the initialize function must match

the GPIB address of the instrument.

- there is one unused device name in the handler

________________________________________________________________________________

LabWindows Libraries:

In order to use this instrument module outside of the LabWindows envi-

ronment, it must be linked with the following LabWindows libraries:

- Formatting and I/O

- GPIB

________________________________________________________________________________

Error Codes:

All error codes are returned as an integer value.

A program should examine this value after each function call is made

to determine if any errors occurred

________________________________________________________________________________

Function Tree Layout:

Kepco SN488 Power Supply Programmer

Initialize

Configure

Set Output

Close

________________________________________________________________________________

Function           : Initialize

QuickBASIC Syntax  : Error% = SN488.init(Address%)

C Syntax           : int SN488_init(Address);

int Address;

Purpose            :  Panel Name:    Initialize

Description:   opens the instrument

and sets its GPIB address

to the specified value

Parameter List     : Address - Integer

Control Name:  Address

Description:   Specifies the address on the

GPIB to be used to communicate

with the instrument.  This value

must correspond to the current

internal setting of the

instrument.

Variable Type: Integer

Valid Range:   0 to 30

Return Value       :   Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Returned Values:

0     Success

-1     GPIB Address out of range

220     Unable to open instrument

223     Device at specified address is not

the correct instrument

230     GPIB write error

233     Unable to configure GPIB address

QuickBASIC Example : Error% = SN488.init(??)

C Example          : int SN488_init(??);

________________________________________________________________________________

Function           : Configure

QuickBASIC Syntax  : Error% = SN488.config(Vrange%, Crange%, _

Programmer.Type%, Maximum.Voltage#, _

Maximum.Current#)

C Syntax           : int SN488_config(Vrange, Crange, Programmer_Type,

Maximum_Voltage, Maximum_Current);

int Vrange;

int Crange;

int Programmer_Type;

double Maximum_Voltage;

double Maximum_Current;

Purpose            : Panel Name : Configure

Description :

Sets up a the SN488 programmer

to be used for a particular

power supply model.

Parameter List     : Vrange - Integer

0 - +10 (Default)

1 - -10

2 - +1

3 - -1

Crange - Integer

0 - +10 (Default)

1 - -10

2 - +1

3 - -1

Programmer Type - Integer

1 - 031,032

0 - 121,122 (Default)

Maximum Voltage - Real

Control Name : Maximum Voltage

Description:

Sets the full scale voltage value.

This value should be set to the maximum

rated voltage value of the power supply

that will be programmed.

Variable Type : Real

Valid Range : 0 - 10000

Maximum Current - Real

Control Name : Maximum Current

Description:

Sets the full scale current value.

This value should be set to the maximum

rated current value of the power supply

that will be programmed.

Variable Type : Real

Valid Range : 0 - 1000

Return Value       : Control name:   Error

Description:   Displays the results of the

function call.

Variable type:  Integer

Returned Values:

0    Success

-1    Invalid voltage range

-2    Invalid current range

-3    Invalid programmer type

-4    Invalid voltage value

-5    Invalid current value

230    GPIB write error

232    Instrument not initialized

QuickBASIC Example : Error% = SN488.config(??, ??, ??, ??, ??)

C Example          : int SN488_config(??, ??, ??, ??, ??);

________________________________________________________________________________

Function           : Set Output

QuickBASIC Syntax  : Error% = SN488.set(Channel%, Voltage#, Current#)

C Syntax           : int SN488_set(Channel, Voltage, Current);

int Channel;

double Voltage;

double Current;

Purpose            : Panel Name : Set Output

Description :

Programs the output voltage and current

of a selected power supply.

Parameter List     : Channel - Integer

0 - Both (Default)

1 - Current

2 - Voltage

Voltage - Real

Control name:   Voltage

Description:

Selects the value of voltage to be programmed.

The full scale value must have been previously

set by the configure function, otherwise a

value of 10 will correspond to full scale.

Variable Type : Real

Valid Range is set by the configure function

Current - Real

Control name:   Current

Description:

Selects the value of current to be programmed.

The full scale value must have been previously

set by the configure function, otherwise a

value of 10 will correspond to full scale.

Variable Type : Real

Valid Range is set by the configure function

Return Value       : Control name:   Error

Description:   Displays the results of the

function call.

Variable type:  Integer

Returned Values:

0    Success

-1    Channel out of range

-2    Voltage value out of range

-3    Current value out of range

230    GPIB write error

232    Instrument not initialized

QuickBASIC Example : Error% = SN488.set(??, ??, ??)

C Example          : int SN488_set(??, ??, ??);

________________________________________________________________________________

Function           : Close

QuickBASIC Syntax  : Error% = SN488.close

C Syntax           : int SN488_close();

Purpose            :  Panel Name:    Close

Description:   Takes the instrument off-line.

Note: The  instrument must  be reinitialized

to use it again.

Return Value       : Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Returned Values:

0     Success

221     Unable to close instrument

232     Instrument not initialized

QuickBASIC Example : Error% = SN488.close

C Example          : int SN488_close();